Point-by-Point Response to Reviewer Comments (Major Revision)

Manuscript: A low-complexity comparative framework for perivascular liver radiofrequency ablation: shape-based assessment of safety-margin deficit and protocol trade-offs

We sincerely thank the reviewer for the careful and constructive evaluation. The comments remained highly valuable in sharpening both the methodological positioning and the transparency of the manuscript. In response, we have revised the manuscript to further reduce overly strong claims, clarified that the moderate-energy protocol is a deliberately conservative comparator rather than a margin-satisfying baseline, added explicit absolute collateral-damage descriptors, inserted measured runtime information obtained on an Apple M4 MacBook Air, expanded the literature-positioning discussion, and further clarified the reduced-order comparative nature of the model.

Comment 1. Validation remains limited.

Response: We agree that the present model does not provide external experimental or clinical validation. Accordingly, we further down-positioned the manuscript as a proof-of-concept methodological study and avoided predictive or clinically deployable wording. We also expanded the related-work discussion and explicitly distinguish the present internal benchmark from external validation. While no new external dataset was added, the revised text now frames the benchmark as literature-consistent comparative evidence rather than predictive validation.

Changes made in manuscript: Abstract, Introduction, Methods (Sections 2.6–2.8), Discussion, Conclusion, and Code/Data Availability statements revised; wording further aligned with a proof-of-concept reduced-order comparative study.

Comment 2. The balanced protocol is confounded because it does not satisfy the no-vessel 5-mm margin target.

Response: We agree and now state this more explicitly. The former “balanced” protocol is renamed the moderate-energy protocol and is described as a deliberately conservative, subtherapeutic comparator rather than as a no-vessel margin-satisfying baseline. We now clarify that the protocol comparison should be interpreted as an analysis of how protocol strength interacts with perivascular cooling, not as a pure compensation analysis around a fully adequate baseline.

Changes made in manuscript: Abstract, Introduction, Methods Section 2.6, Results Section 3.4, Discussion, Conclusion, and terminology throughout manuscript and captions revised.

Comment 3. The protocol effect size is close to discretization-induced variation.

Response: We agree that this limits the strength of the inference. We therefore further softened the interpretation of MDI. The revised text now states that MDI is most appropriate as a comparative ranking and trend descriptor within the tested discretization range, rather than as a high-precision absolute estimate of clinical margin failure.

Changes made in manuscript: Abstract (Results sentence), Results Section 3.5, and Discussion revised.

Comment 4. The model is a strongly simplified reduced-order proxy and its scope should be made explicit.

Response: We agree. The revised manuscript now consistently describes the framework as a reduced-order comparative model or planning-oriented comparative proxy rather than a predictive patient-specific simulator. We additionally clarify that the credibility chain is limited to single-point no-vessel calibration, an internal normalized benchmark, and literature-consistent trends, which should not be interpreted as external validation.

Changes made in manuscript: Title, Abstract, Methods (Sections 2.2-2.8), Discussion, and Conclusion revised.

Comment 5. The term “planning model” is too broad.

Response: We agree. To avoid overstating the scope, we replaced broad “planning model” wording with “planning-oriented comparative framework”, “planning-oriented comparative analysis”, or equivalent lower-claim wording throughout the manuscript.

Changes made in manuscript: Title, Introduction, Results, Discussion, and Conclusion revised.

Comment 6. The literature positioning and innovation framing were too limited.

Response: We agree. The literature context has been expanded to include prior work on patient-specific simulation workflows, prospective in silico planning, fusion-imaging based assessment, and margin quantification. We now explicitly position the study as an incremental but useful methodological simplification rather than as a fundamentally new physical model.

Changes made in manuscript: Introduction and Discussion revised; reference list expanded.

Comment 7. The trade-off analysis needs a more direct measure of collateral damage.

Response: We agree. To complement PPV_target, we added supplementary analyses reporting (A) absolute overablation area outside the target region and (B) target-region coverage. These additions make the trade-off analysis less dependent on proportional specificity alone and provide a more direct description of collateral damage.

Changes made in manuscript: Methods Section 2.4 revised; new Fig. S7 added; Results Section 3.4 and Discussion revised.

Comment 8. The protocol space is sparse and conclusions should be restricted to the tested range.

Response: We agree and have made this limitation explicit. The revised manuscript now repeatedly states that protocol-dependent conclusions hold only within the tested protocol and discretization ranges, and should not be interpreted as a general protocol landscape or optimization result.

Changes made in manuscript: Methods Section 2.6; Results Section 3.3; Discussion revised.

Comment 9. The claim of computational lightness requires measured runtime.

Response: We agree. We measured runtime on an Apple M4 MacBook Air. On the final 161 × 161 grid, a representative single-case run required approximately 2.4 s, while the final 18-case protocol sweep required approximately 13.3 s. Peak memory footprints were approximately 190 MB and 282 MB, respectively. These values are now reported in the manuscript to support the low-complexity characterization.

Changes made in manuscript: Methods Section 2.8 revised with wall-clock time and memory information; wording in Abstract and Discussion kept aligned with a low-complexity comparative framework.

We hope that these revisions address the reviewer’s concerns and further clarify both the intended scope and the methodological contribution of the work. The manuscript is now framed consistently as a reduced-order, planning-oriented comparative framework for proof-of-concept analysis under idealized perivascular conditions.